## Sheet1
Mult. | 1.219 | For bb
Sample #: | 1 |  |  | Sample #: | 2
Catheter: | B04002-RD |  |  | Catheter: | B04003-RD
| -------------VNA Test Fully Assembled Catheter----------------- |  |  | -------------VNA Test Fully Assembled Catheter-----------------
Channel | Capacitance @ Low Frequency (pF) | Short/Open? | Pass/Fail | Channel | Capacitance @ Low Frequency (pF) | Short/Open? | Pass/Fail
1 | 787 |  | Pass | 1 | 785 |  | Pass
2 | 744 |  | Pass | 2 | 764 |  | Pass
3 | 762 |  | Pass | 3 | 781 |  | Pass
4 | 737 |  | Pass | 4 | 407 |  | Fail
5 | 745 |  | Pass | 5 | 753 |  | Pass
6 | 756 |  | Pass | 6 | 351 |  | Fail
7 | 726 |  | Pass | 7 | 402 |  | Fail
8 | 777 |  | Pass | 8 | 787 |  | Pass
9 | 747 |  | Pass | 9 | 772 |  | Pass
10 | 435 |  | Fail | 10 | 821 |  | Fail
11 | 732 |  | Pass | 11 | 742 |  | Pass
12 | 788 |  | Pass | 12 | 797 |  | Pass
13 | 757 |  | Pass | 13 | 757 |  | Pass
14 | 756 |  | Pass | 14 | 766 |  | Pass
15 | 746 |  | Pass | 15 | 757 |  | Pass
16 | 736 |  | Pass | 16 | 403 |  | Fail
17 | 745 |  | Pass | 17 | 756 |  | Pass
18 | 735 |  | Pass | 18 | 756 |  | Pass
19 | 738 |  | Pass | 19 | 760 |  | Pass
20 | 594 |  | Fail | 20 | 761 |  | Pass
21 | 775 |  | Pass | 21 | 801 |  | Pass
22 | 756 |  | Pass | 22 | 27uh |  | Fail
23 | 713 |  | Pass | 23 | 750 |  | Pass
24 | 743 |  | Pass | 24 | 754 |  | Pass
25 | 742 |  | Pass | 25 | 741 |  | Pass
26 | 190 | Open | Fail | 26 | 767 |  | Pass
27 | 771 |  | Pass | 27 | 776 |  | Pass
28 | 787 |  | Pass | 28 | 815 |  | Pass
29 | 742 |  | Pass | 29 | 25uh |  | Fail
30 | 750 |  | Pass | 30 | 753 |  | Pass
31 | 737 |  | Pass | 31 | 746 |  | Pass
32 | 749 |  | Pass | 32 | 760 |  | Pass
33 | 733 |  | Pass | 33 | 743 |  | Pass
34 | 729 |  | Pass | 34 | 742 |  | Pass
35 | 711 |  | Pass | 35 | 627 |  | Fail
36 | 784 |  | Pass | 36 | 445 |  | Fail
37 | 748 |  | Pass | 37 | 624 |  | Fail
38 | 793 |  | Pass | 38 | 815 |  | Pass
39 | 737 |  | Pass | 39 | 746 |  | Pass
40 | 736 |  | Pass | 40 | 752 |  | Pass
41 | 749 |  | Pass | 41 | 770 |  | Pass
42 | 785 |  | Pass | 42 | 802 |  | Pass
43 | 765 |  | Pass | 43 | 785 |  | Pass
44 | 714 |  | Pass | 44 | 770 |  | Pass
45 | 710 |  | Pass | 45 | 759 |  | Pass
46 | 793 |  | Pass | 46 | 805 |  | Pass
47 | 740 |  | Pass | 47 | 757 |  | Pass
48 | 767 |  | Pass | 48 | 780 |  | Pass
49 | 739 |  | Pass | 49 | 764 |  | Pass
50 | 748 |  | Pass | 50 | 764 |  | Pass
51 | 627 |  | Fail | 51 | 624 |  | Fail
52 | 753 |  | Pass | 52 | 772 |  | Pass
53 | 627 |  | Fail | 53 | 761 |  | Pass
54 | 746 |  | Pass | 54 | 761 |  | Pass
55 | 721 |  | Pass | 55 | 766 |  | Pass
56 | 752 |  | Pass | 56 | 773 |  | Pass
57 | 752 |  | Pass | 57 | 752 |  | Pass
58 | 752 |  | Pass | 58 | 775 |  | Pass
59 | 802 |  | Pass | 59 | 810 |  | Pass
60 | 708 |  | Pass | 60 | 764 |  | Pass
61 | 740 |  | Pass | 61 | 754 |  | Pass
62 | 744 |  | Pass | 62 | 765 |  | Pass
63 | 750 |  | Pass | 63 | 771 |  | Pass
64 | 787 |  | Pass | 64 | 813 |  | Pass
Sample #: | 3 |  |  | Sample #: | 4
Catheter: | B04007-RD |  |  | Catheter: | B04009-RD
| -------------VNA Test Fully Assembled Catheter----------------- |  |  |  | -------------VNA Test Fully Assembled Catheter-----------------
Channel | Capacitance @ Low Frequency (pF) | Short/Open? | Pass/Fail | Channel | Capacitance @ Low Frequency (pF) | Short/Open? | Pass/Fail
1 | 658 |  | Fail | 1 | 839 |  | Fail
2 | 752 |  | Pass | 2 | 797 |  | Pass
3 | 1112 |  | Fail | 3 | 820 |  | Pass
4 | 644 |  | Fail | 4 | 794 |  | Pass
5 | 728 |  | Pass | 5 | 807 |  | Pass
6 | 1094 |  | Fail | 6 | 187 |  | Fail
7 | 723 |  | Pass | 7 | 790 |  | Pass
8 | 781 |  | Pass | 8 | 824 |  | Fail
9 | 1088 |  | Fail | 9 | 187 |  | Fail
10 | 1154 |  | Fail | 10 | 189 |  | Fail
11 | 735 |  | Pass | 11 | 786 |  | Pass
12 | 797 |  | Pass | 12 | 828 |  | Fail
13 | 776 |  | Pass | 13 | 776 |  | Pass
14 | 768 |  | Pass | 14 | 814 |  | Pass
15 | 749 |  | Pass | 15 | 802 |  | Pass
16 | 740 |  | Pass | 16 | 798 |  | Pass
17 | 758 |  | Pass | 17 | 807 |  | Pass
18 | 745 |  | Pass | 18 | 790 |  | Pass
19 | 748 |  | Pass | 19 | 795 |  | Pass
20 | 760 |  | Pass | 20 | 801 |  | Pass
21 | 787 |  | Pass | 21 | 822 |  | Fail
22 | 510 |  | Fail | 22 | 821 |  | Fail
23 | 414 |  | Fail | 23 | 800 |  | Pass
24 | 756 |  | Pass | 24 | 805 |  | Pass
25 | 744 |  | Pass | 25 | 798 |  | Pass
26 | 1101 |  | Fail | 26 | 994 |  | Fail
27 | 670 |  | Pass | 27 | 815 |  | Pass
28 | 790 |  | Pass | 28 | 829 |  | Fail
29 | 750 |  | Pass | 29 | 790 |  | Pass
30 | 754 |  | Pass | 30 | 809 |  | Pass
31 | 748 |  | Pass | 31 | 799 |  | Pass
32 | 57uh |  | Fail | 32 | 828 |  | Fail
33 | 740 |  | Pass | 33 | 788 |  | Pass
34 | 731 |  | Pass | 34 | 786 |  | Pass
35 | 895 |  | Fail | 35 | 678 |  | Pass
36 | 792 |  | Pass | 36 | 815 |  | Pass
37 | 755 |  | Pass | 37 | 796 |  | Pass
38 | 801 |  | Pass | 38 | 824 |  | Fail
39 | 743 |  | Pass | 39 | 794 |  | Pass
40 | 416 |  | Fail | 40 | 798 |  | Pass
41 | 743 |  | Pass | 41 | 806 |  | Pass
42 | 782 |  | Pass | 42 | 826 |  | Fail
43 | 768 |  | Pass | 43 | 824 |  | Fail
44 | 1083 |  | Fail | 44 | 184 | open | Fail
45 | 743 |  | Pass | 45 | 802 |  | Pass
46 | 802 |  | Pass | 46 | 840 |  | Fail
47 | 740 |  | Pass | 47 | 803 |  | Pass
48 | 776 |  | Pass | 48 | 808 |  | Pass
49 | 747 |  | Pass | 49 | 802 |  | Pass
50 | 742 |  | Pass | 50 | 804 |  | Pass
51 | 757 |  | Pass | 51 | 776 |  | Pass
52 | 750 |  | Pass | 52 | 815 |  | Pass
53 | 776 |  | Pass | 53 | 776 |  | Pass
54 | 1082 |  | Fail | 54 | 188 | open | Fail
55 | 750 |  | Pass | 55 | 796 |  | Pass
56 | 757 |  | Pass | 56 | 798 |  | Pass
57 | 757 |  | Pass | 57 | 776 |  | Pass
58 | 1081 |  | Fail | 58 | 831 |  | Fail
59 | 807 |  | Pass | 59 | 837 |  | Fail
60 | 745 |  | Pass | 60 | 805 |  | Pass
61 | 745 |  | Pass | 61 | 797 |  | Pass
62 | 744 |  | Pass | 62 | 808 |  | Pass
63 | 743 |  | Pass | 63 | 815 |  | Pass
64 | 791 |  | Pass | 64 | 835 |  | Fail
Sample #: | 5 |  |  | Sample #: | 6
Catheter: | B04010-RD |  |  | Catheter: | B04011-RD
| -------------VNA Test Fully Assembled Catheter----------------- |  |  |  | -------------VNA Test Fully Assembled Catheter-----------------
Channel | Capacitance @ Low Frequency (pF) | Short/Open? | Pass/Fail | Channel | Capacitance @ Low Frequency (pF) | Short/Open? | Pass/Fail
1 | 781 |  | Pass | 1 | 437 |  | Fail
2 | 740 |  | Pass | 2 | 411 |  | Fail
3 | 756 |  | Pass | 3 | 351 |  | Fail
4 | 733 |  | Pass | 4 | 400 |  | Fail
5 | 750 |  | Pass | 5 | 680 |  | Pass
6 | 746 |  | Pass | 6 | 336 |  | Fail
7 | 690 |  | Pass | 7 | 398 |  | Fail
8 | 767 |  | Pass | 8 | 435 |  | Fail
9 | 744 |  | Pass | 9 | 331 |  | Fail
10 | 782 |  | Pass | 10 | 375 |  | Fail
11 | 726 |  | Pass | 11 | 390 |  | Fail
12 | 773 |  | Pass | 12 | 443 |  | Fail
13 | 776 |  | Pass | 13 | 776 |  | Pass
14 | 760 |  | Pass | 14 | 412 |  | Fail
15 | 711 |  | Pass | 15 | 409 |  | Fail
16 | 727 |  | Pass | 16 | 394 |  | Fail
17 | 749 |  | Pass | 17 | 750 |  | Pass
18 | 732 |  | Pass | 18 | 397 |  | Fail
19 | 735 |  | Pass | 19 | 400 |  | Fail
20 | 742 |  | Pass | 20 | 412 |  | Fail
21 | 764 |  | Pass | 21 | 430 |  | Fail
22 | 750 |  | Pass | 22 | 344 |  | Fail
23 | 742 |  | Pass | 23 | 398 |  | Fail
24 | 749 |  | Pass | 24 | 405 |  | Fail
25 | 739 |  | Pass | 25 | 395 |  | Fail
26 | 760 |  | Pass | 26 | 760 |  | Pass
27 | 768 |  | Pass | 27 | 420 |  | Fail
28 | 772 |  | Pass | 28 | 437 |  | Fail
29 | 737 |  | Pass | 29 | 407 |  | Fail
30 | 753 |  | Pass | 30 | 404 |  | Fail
31 | 745 |  | Pass | 31 | 399 |  | Fail
32 | 755 |  | Pass | 32 | 272 |  | Fail
33 | 726 |  | Pass | 33 | 392 |  | Fail
34 | 724 |  | Pass | 34 | 391 |  | Fail
35 | 623 |  | Fail | 35 | 278 |  | Fail
36 | 769 |  | Pass | 36 | 441 |  | Fail
37 | 740 |  | Pass | 37 | 270bb |  | Fail
38 | 774 |  | Pass | 38 | 443 |  | Fail
39 | 733 |  | Pass | 39 | 402 |  | Fail
40 | 744 |  | Pass | 40 | 398 |  | Fail
41 | 757 |  | Pass | 41 | 416 |  | Fail
42 | 787 |  | Pass | 42 | 443 |  | Fail
43 | 772 |  | Pass | 43 | 428 |  | Fail
44 | 738 |  | Pass | 44 | 330 |  | Fail
45 | 745 |  | Pass | 45 | 746 |  | Pass
46 | 785 |  | Pass | 46 | 443 |  | Fail
47 | 743 |  | Pass | 47 | 404 |  | Fail
48 | 758 |  | Pass | 48 | 424 |  | Fail
49 | 747 |  | Pass | 49 | 757 |  | Pass
50 | 750 |  | Pass | 50 | 756 |  | Pass
51 | 776 |  | Pass | 51 | 776 |  | Pass
52 | 756 |  | Pass | 52 | 414 |  | Fail
53 | 776 |  | Pass | 53 | 776 |  | Pass
54 | 742 |  | Pass | 54 | 325 |  | Fail
55 | 744 |  | Pass | 55 | 410 |  | Fail
56 | 759 |  | Pass | 56 | 764 |  | Pass
57 | 776 |  | Pass | 57 | 776 |  | Pass
58 | 760 |  | Pass | 58 | 776 |  | Pass
59 | 792 |  | Pass | 59 | 793 |  | Pass
60 | 748 |  | Pass | 60 | 405 |  | Fail
61 | 735 |  | Pass | 61 | 401 |  | Fail
62 | 749 |  | Pass | 62 | 407 |  | Fail
63 | 759 |  | Pass | 63 | 414 |  | Fail
64 | 794 |  | Pass | 64 | 451 |  | Fail